## Slide 1: PRM Method Development and Data Analysis with Skyline
Tutorial Webinar #17
With
Brendan MacLean (Principal Developer, Skyline)
Eduard Sabidó, Ph.D. (Head of the UPF/CRG Proteomics Unit)
Cristina Chiva, Ph.D. (PRM researcher, CRG Proteomics Unit)

## Slide 2: Agenda
Welcome from the Skyline team!
PRM Method Development and Data Analysis with Skyline
  Introduction with Brendan MacLean
  Theoretical concepts and benefits of PRM with Eduard Sabidó
  Tutorial with Cristina Chiva
Audience Q&A – submit questions to Google Form:
https://skyline.ms/QA4Skyline.url

## Slide 3: It Began as Targeted MS/MS (pseudo-SRM)
ASMS 2011 – Poster presentation (Birgit Schilling)
  Skyline: Targeted Proteomics with Extracted Ion Chromatograms from Full-Scan Mass Spectra
JPR May, 2012 – Sherrod, et. al (started by Amy Ham, 2009)
  Label-Free Quantitation of Protein Modifications by Pseudo-Selected Reaction Monitoring with Internal Reference Peptides
MCP Nov, 2012 – Peterson, et. al (Coon lab) – named PRM
  Parallel reaction monitoring for high resolution and high mass accuracy quantitative, targeted proteomics
Anal. Chem. 2015 – Schilling, et. al – latest on Skyline PRM for HRMS
  Multiplexed, Scheduled, High-Resolution Parallel Reaction Monitoring on a Full Scan QqTOF…

## Slide 4: Recorded Webinars
16 other webinars + this one Coming Soon!
Webinar #3 (2013) PRM Targeted Proteomics Using Full-Scan MS and Skyline Bruno Domon
Webinar #9 (2015) PRM for PTM studies with Skyline Research-grade targeted proteomics assay development Jacob D. Jaffe

## Slide 5: Interest Remains High
Modern instruments support PRM well
  High resolution, faster cycle times and full scheduling
Viable alternative to SRM with triple quadrupole instrument
Appealing not to need an extra instrument for targeted
Registration for first PRM webinar (431) v 2014 DIA webinar (342)
Registration for this PRM webinar (434) v 2017 DIA webinar (409)

## Slide 6: Eduard Sabidó

## Slide 7
Webinar
PRM theoretical concepts, benefits and instrument acquisition settings
Eduard Sabidó
Cristina Chiva
CRG/UPF Proteomics Unit
Barcelona, Spain
Notes: 7

## Slide 8
Lecture on PRM theoretical concepts, benefits and instrument acquisition settings for the Orbitrap Fusion Lumos (30 min).
Hands-on tutorial on PRM method development and step-by-step analysis of Orbitrap Fusion Lumos PRM data with Skyline (1h)
Questions & Answers
Eduard Sabidó
Cristina Chiva

## Slide 9
PRM is a targeted proteomics workflow
“Targeted proteomics detects proteins of interest with high sensitivity, quantitative accuracy and reproducibility”
“By delivering precise, reproducible quantification of proteins of interest in biological samples, targeted proteomics approaches are allowing researchers to apply the scientific method using mass spectrometry”

## Slide 10
PRM is a targeted proteomics workflow
Types of projects suited for targeted proteomics
B pathways
A biomarkers
C interactions

## Slide 11
PRM is a targeted proteomics workflow
MS2 Targeted Methods
They rely on the
fragments of the molecule
MS1 Targeted Methods
They rely on the
mass of the entire molecule
Targeted Data Analysis
They acquire everything and later
specific information is extracted
Targeted Acquisition
They only acquire the
molecules of interest

## Slide 12
PRM is a targeted proteomics workflow
Parallel Reaction Monitoring (PRM)
Gallien S, ..., Domon B, Mol Cell Proteomics. 2012 Dec; 11(12): 1709-1723
Peterson, ..., Coon J, Mol Cell Proteomics. 2012 Nov; 11(11): 1475-1488

## Slide 13
PRM is a targeted proteomics workflow
Parallel Reaction Monitoring (PRM)
Gallien S, ..., Domon B, Mol Cell Proteomics. 2012 Dec; 11(12): 1709-1723
Peterson, ..., Coon J, Mol Cell Proteomics. 2012 Nov; 11(11): 1475-1488

## Slide 14
PRM is a targeted proteomics workflow
Parallel Reaction Monitoring (PRM)
co-elution
identification

## Slide 15
PRM is a targeted proteomics workflow
Selected Reaction Monitoring (SRM) is sequential (not parallel)

## Slide 16
MS1
MS2
TOF / OT
QUADRUPOLE
Low-resolution MS2 signal
One narrow scan per fragment
Hardware selected MS2 fragments
m/z
(m/z)1
(m/z)2
(m/z)3
QUADRUPOLE
QUADRUPOLE
SRM
PRM
PRM is a targeted proteomics workflow
Mass spectrometry instruments
High-resolution MS2 signal
Full scan MS2
Software selected MS2 fragments
Isolation window 0.7 m/z
Isolation window 1.4 m/z
Isolation window 0.7 m/z

## Slide 17
1
2
3
PRM is a targeted proteomics workflow
PRM in the Orbitrap Fusion Lumos
Precursor selection in quadrupole
    Isolation window
Precursor fragmentation in collision cell
    Fill time
    Collision energy
Fragment ion detection in the Orbitrap
    Resolving power
Notes: In PRM mode, a limited m/z range including the precursor ions of interest is isolated by the quadrupole (1), transmitted via the C-trap to the HCD cell where corresponding fragment ions are generated and accumulated (2). Fragment ions are then transferred back into the C-trap (3) and eventually injected and analyzed in the orbitrap mass analyzer (4). The actual quantification of endogenous peptides is performed on the peak areas of corresponding selected fragment ions.
17

## Slide 18
Peptide verification
Making sure that the integrated signal corresponds to the targeted peptide

## Slide 19
Peptide verification
Making sure that the integrated signal corresponds to the targeted peptide
1. Co-elution of concurrent transitions

## Slide 20
Peptide verification
Making sure that the integrated signal corresponds to the targeted peptide
1. Co-elution of concurrent transitions
2. Sequence information / coverage

## Slide 21
Peptide verification
Making sure that the integrated signal corresponds to the targeted peptide
1. Co-elution of concurrent transitions
2. Sequence information / coverage
3. Reference MS2 spectra (libraries)

## Slide 22
Peptide verification
Making sure that the integrated signal corresponds to the targeted peptide

## Slide 23
Peptide verification
Making sure that the integrated signal corresponds to the targeted peptide
1. Co-elution of concurrent transitions
2. Sequence information / coverage
3. Reference MS2 spectra (libraries)
4. Retention time information

## Slide 24
Peptide verification
Making sure that the integrated signal corresponds to the targeted peptide

## Slide 25
Peptide verification
Making sure that the integrated signal corresponds to the targeted peptide

## Slide 26
Peptide verification
Making sure that the integrated signal corresponds to the targeted peptide
Notes: Reduce matrix complexity: fractionation?
Low throughput
Quantitative challenge (reproducibility, robustness)
26

## Slide 27
Peptide verification
Making sure that the integrated signal corresponds to the targeted peptide
1. Co-elution of concurrent transitions
2. Sequence information / coverage
3. Reference MS2 spectra (libraries)
4. Retention time information
5. Reference internal standard

## Slide 28
Peptide verification
Making sure that the integrated signal corresponds to the targeted peptide

## Slide 29
Peptide quantification
Aim for a peptide quantification with high accuracy and precision

## Slide 30
1
2
3
4
5
6
7
8
9
10
8-10 points across chromatographic peak to define elution profile
Peptide quantification
Aim for a peptide quantification with high accuracy and precision
Notes: Importance of experimental design
30

## Slide 31
4 points across chromatography peak
Theoretical
Healthy
Disease
Theoretical
Healthy
Disease
Peptide quantification
Aim for a peptide quantification with high accuracy and precision

## Slide 32
Peptide quantification
Calibration strategies with internal standards
Single-point calibration
External calibration curve
Reverse calibration curve
Aim for a peptide quantification with high accuracy and precision
Webinar #13 (2016)
Calibrated Quantification with Skyline
Chris Shuford
Notes: Importance of experimental design
32

## Slide 33
Peptide quantification
Aim for a peptide quantification with high accuracy and precision
Area
Concentration
50
100
150
200
250
300
We establish the calibration line with two points
  The response of the heavy internal standard
  The zero — that is, zero response for zero concentration.
We assume a linear range response
0
0
heavy internal standard
endogenous
peptide
Single-point calibration
Notes: Importance of experimental design
33

## Slide 34
Peptide quantification
Aim for a peptide quantification with high accuracy and precision
Area
Concentration
50
100
150
200
250
300
0
0
heavy internal standard
endogenous
peptide
LOD
LLOQ
ULOQ
Single-point calibration
We establish the calibration line with two points
  The response of the heavy internal standard
  The zero — that is, zero response for zero concentration.
We assume a linear range response
Notes: Importance of experimental design
34

## Slide 35
1
2
3
Precursor selection in quadrupole
    Isolation window
Precursor fragmentation in collision cell
    Fill time
    Collision energy
Fragment ion detection in the Orbitrap
    Resolving power
Sensitivity and selectivity: a double boost for PRM
PRM in the Orbitrap Fusion Lumos
Notes: In PRM mode, a limited m/z range including the precursor ions of interest is isolated by the quadrupole (1), transmitted via the C-trap to the HCD cell where corresponding fragment ions are generated and accumulated (2). Fragment ions are then transferred back into the C-trap (3) and eventually injected and analyzed in the orbitrap mass analyzer (4). The actual quantification of endogenous peptides is performed on the peak areas of corresponding selected fragment ions.
35

## Slide 36
1
2
3
Sensitivity and selectivity: a double boost for PRM
PRM in the Orbitrap Fusion Lumos
Precursor selection in quadrupole
    Isolation window
Precursor fragmentation in collision cell
    Fill time
    Collision energy
Fragment ion detection in the Orbitrap
    Resolving power
SELECTIVITY
SENSITIVITY
Notes: In PRM mode, a limited m/z range including the precursor ions of interest is isolated by the quadrupole (1), transmitted via the C-trap to the HCD cell where corresponding fragment ions are generated and accumulated (2). Fragment ions are then transferred back into the C-trap (3) and eventually injected and analyzed in the orbitrap mass analyzer (4). The actual quantification of endogenous peptides is performed on the peak areas of corresponding selected fragment ions.
36

## Slide 37
Fill time = 60 ms Res Power = 30K
Sensitivity and selectivity: a double boost for PRM
PRM in the Orbitrap Fusion Lumos
Notes: In PRM mode, a limited m/z range including the precursor ions of interest is isolated by the quadrupole (1), transmitted via the C-trap to the HCD cell where corresponding fragment ions are generated and accumulated (2). Fragment ions are then transferred back into the C-trap (3) and eventually injected and analyzed in the orbitrap mass analyzer (4). The actual quantification of endogenous peptides is performed on the peak areas of corresponding selected fragment ions.
37

## Slide 38
Fill time = 250 ms Res Power = 120K
Sensitivity and selectivity: a double boost for PRM
PRM in the Orbitrap Fusion Lumos
Notes: In PRM mode, a limited m/z range including the precursor ions of interest is isolated by the quadrupole (1), transmitted via the C-trap to the HCD cell where corresponding fragment ions are generated and accumulated (2). Fragment ions are then transferred back into the C-trap (3) and eventually injected and analyzed in the orbitrap mass analyzer (4). The actual quantification of endogenous peptides is performed on the peak areas of corresponding selected fragment ions.
38

## Slide 39
1
2
3
P1
P2
P3
P4
P5
120ms
Fill Time (120 ms)
Transient length OT (60k 120ms)
Sensitivity and selectivity: a double boost for PRM
PRM in the Orbitrap Fusion Lumos
Notes: In PRM mode, a limited m/z range including the precursor ions of interest is isolated by the quadrupole (1), transmitted via the C-trap to the HCD cell where corresponding fragment ions are generated and accumulated (2). Fragment ions are then transferred back into the C-trap (3) and eventually injected and analyzed in the orbitrap mass analyzer (4). The actual quantification of endogenous peptides is performed on the peak areas of corresponding selected fragment ions.
39

## Slide 40
Fill Time (120 ms)
P1
P2
P3
P4
P5
120ms
P1
P2
P3
P4
P5
120 ms
P1
P2
P3
P4
P5
120ms
Fill Time (60 ms)
Fill Time (120 ms)
Transient length OT (30k 60ms)
1
2
3
Sensitivity and selectivity: a double boost for PRM
PRM in the Orbitrap Fusion Lumos
Transient length OT (60k 120ms)
Transient length OT (60k 120ms)
Notes: In PRM mode, a limited m/z range including the precursor ions of interest is isolated by the quadrupole (1), transmitted via the C-trap to the HCD cell where corresponding fragment ions are generated and accumulated (2). Fragment ions are then transferred back into the C-trap (3) and eventually injected and analyzed in the orbitrap mass analyzer (4). The actual quantification of endogenous peptides is performed on the peak areas of corresponding selected fragment ions.
40

## Slide 41
Sensitivity and selectivity: a double boost for PRM
PRM in the Orbitrap Fusion Lumos

## Slide 42
THROUGHPUT
# peptides
SELECTIVITY
OT Resolution
SENSITIVITY
Fill Time
CYCLE
TIME
Sensitivity and selectivity: a double boost for PRM
PRM in the Orbitrap Fusion Lumos

## Slide 43
THROUGHPUT
# peptides
SELECTIVITY
OT Resolution
SENSITIVITY
Fill Time
CYCLE
TIME
Sensitivity and selectivity: a double boost for PRM
PRM in the Orbitrap Fusion Lumos
Scheduled PRM
Chromatographic gradient
time
0
90
Peptide 1
Peptide 2
Peptide 3
Peptide 4
Peptide 5

## Slide 44
Orbitrap Fusion Lumos Method Editor for PRM

## Slide 46
Orbitrap Fusion Lumos Method Editor for PRM

## Slide 47
Webinar #3 (2013)
PRM Targeted Proteomics Using Full-Scan MS and Skyline
Bruno Domon
Webinar #9 (2015)
Research-grade targeted proteomics assay development: PRM for PTM studies with Skyline.
Jacob D. Jaffe
Other webinars
https://skyline.ms/wiki/home/software/Skyline/page.view?name=webinars
More PRM Webinars

## Slide 48
Hands-on tutorial
PRM method development and step-by-step analysis of Lumos PRM data

## Slide 49
31 peptides from 19 proteins related to the cell cycle of mouse fibroblasts
Hands-on tutorial
PRM method development and step-by-step analysis of Lumos PRM data

## Slide 50
1. Setting up a Skyline document for PRM acquisition
2. Prepare and export the PRM method
3. PRM data analysis
Hands-on tutorial
PRM method development and step-by-step analysis of Lumos PRM data

## Slide 51
EMBO Practical Course on Targeted Proteomics
Barcelona, 11-16 November 2018
Registration Opens March 2018
Other courses
https://skyline.ms/project/home/software/Skyline/events/begin.view?
Hands-on tutorial
PRM method development and step-by-step analysis of Lumos PRM data

## Slide 52: Learn More
Webinar #18: TBD (coming soon)
  Listings updated in Join Us section of Skyline homepage:
  https://skyline.ms/Skyline.url
Weeklong Courses 2018
  Buck Institute, Novato, CA – April 2-6
  NEU, Boston – April 30 – May 11
  ETH, Zurich – July 2-6
  U. of Wa., Seattle – July 30 – August 3
  New! Duke, Durham, NC – September 17-21
  CRG, Barcelona – November 12-16
  Shanghai – October 22-26
Workshops and Conferences 2018
  MSACL, Palm Springs – January 20&21
  pre-Lorne, Melbourne – January 29-31
  IIT Bombay, Mumbai – February
  US HUPO, Minneapolis – March 10&11
  ASMS, San Diego – June 2&3
  User Group Meeting at ASMS, San Diego – June 3
  CNPEM, Campinas, Brazil – November 7-9

## Slide 53: Questions?
Ask any questions at the following form:
  https://skyline.ms/QA4Skyline.url
Take the post-webinar survey:
https://skyline.ms/survey4webinar.url

## Slide 54: Tutorial Webinar #17
This ends this Skyline Tutorial Webinar.
Please give us feedback on the webinar at the following survey:
https://skyline.ms/survey4webinar.url
A recording of today’s meeting will be available shortly at the Skyline website.
We look forward to seeing you at a future Skyline Tutorial Webinar.