## Slide 1: Multiplexed Data Independent Acquisition for Comparative Proteomics
Jarrett Egertson
MacCoss Lab
Department of Genome Sciences
University of Washington
5/20/2012

## Slide 2: Current Technology for Comparative Proteomics
Targeted:
  How much does protein X increase/decrease?
  For a small target list (<100 peptides)
  Often requires extra steps
    Retention time scheduling
    Peptide transition refinement
Discovery:
  What proteins are changing in abundance?
  For ~1,000 - 5,000 semi-randomly selected peptides
  Data is not collected on the majority of peptides!

## Slide 3: Many Peptides Are Missed By Data Dependent Acquisition
~25,000 – 50,000 Peptides
Detected in MS
~1,000 – 5,000 Peptides Assigned Sequence Determined By MS/MS
Notes: Usually 1,000 – 5,000 peptides identified by MS/MS
Usually ~25,000 – 50,000 MS features detected
2-20% of Features Annotated
BUT Usually only ~70% overlap in annotated features between runs

## Slide 4: Data Independent Acquisition (DIA) to Increase Sequence Coverage
Scan 1
40 10 m/z-wide windows = 400 m/z
m/z
500
900
Scan 2
Venable JD et. al. Nature Methods 2004.
Notes: Usually 1,000 – 5,000 peptides identified by MS/MS
Usually ~25,000 – 50,000 MS features detected
2-20% of Features Annotated
BUT Usually only ~70% overlap in annotated features between runs

## Slide 5: Data Independent Acquisition (DIA) to Increase Sequence Coverage
Scan 1
40 10 m/z-wide windows = 400 m/z
m/z
500
900
Scan 2
Scan 3
Scan 4
Scan 5
Scan 6
Scan 7
Scan 40
…
Scan 41
Notes: Usually 1,000 – 5,000 peptides identified by MS/MS
Usually ~25,000 – 50,000 MS features detected
2-20% of Features Annotated
BUT Usually only ~70% overlap in annotated features between runs

## Slide 6: Data Independent Acquisition (DIA) to Increase Sequence Coverage
40 10 m/z-wide windows = 400 m/z
m/z
500
900
Retention Time
Notes: Usually 1,000 – 5,000 peptides identified by MS/MS
Usually ~25,000 – 50,000 MS features detected
2-20% of Features Annotated
BUT Usually only ~70% overlap in annotated features between runs

## Slide 7: Targeted-Style Analysis
48
49
50
51
52
Retention Time
0.0
0.5
1.0
1.5
2.0
2.5
3.0
3.5
Intensity x 10-6
LGLVGGSTIDIK++ (586.85)
VGGSTIDIK+
GGSTIDIK+
GSTIDIK+
LVGGSTIDIK+
STIDIK+
TIDIK+
IDIK+
(1002.58)
(889.50)
(790.43)
(676.39)
(589.36)
(488.31)
(375.22)
Notes: Usually 1,000 – 5,000 peptides identified by MS/MS
Usually ~25,000 – 50,000 MS features detected
2-20% of Features Annotated
BUT Usually only ~70% overlap in annotated features between runs

## Slide 8: DIA Lacks the Specificity of DDA
10 m/z
2 m/z
Notes: 10 m/z windows 400 – 800 m/z assuming 10 hz scan frequency

## Slide 9: DIA Interference/Low Specificity
FEIELLSLDDDSIVNHEQDLPK S. cerevisiae lysate (soluble) 10 m/z wide window DIA (Q-Exactive)
Notes: Usually 1,000 – 5,000 peptides identified by MS/MS
Usually ~25,000 – 50,000 MS features detected
2-20% of Features Annotated
BUT Usually only ~70% overlap in annotated features between runs

## Slide 10: Multiplexed DIA
Scan 1
100 4 m/z-wide windows = 400 m/z
m/z
500
900
Notes: Bottom: DIA: 10 m/z wide windows 400 -800 10 hz scan speed
Top: MSX 5 x 2 m/z windows per scan 400 – 800 m/z 10 hz scan speed

## Slide 11: Multiplexed DIA
Scan 1
100 4 m/z-wide windows = 400 m/z
m/z
500
900
Notes: Bottom: DIA: 10 m/z wide windows 400 -800 10 hz scan speed
Top: MSX 5 x 2 m/z windows per scan 400 – 800 m/z 10 hz scan speed

## Slide 12: Multiplexed DIA
Scan 1
100 4 m/z-wide windows = 400 m/z
m/z
500
900
Scan 2
Notes: Bottom: DIA: 10 m/z wide windows 400 -800 10 hz scan speed
Top: MSX 5 x 2 m/z windows per scan 400 – 800 m/z 10 hz scan speed

## Slide 13: Multiplexed DIA
Scan 1
100 4 m/z-wide windows = 400 m/z
m/z
500
900
Scan 2
Notes: Bottom: DIA: 10 m/z wide windows 400 -800 10 hz scan speed
Top: MSX 5 x 2 m/z windows per scan 400 – 800 m/z 10 hz scan speed

## Slide 14: Multiplexed DIA
Scan 1
100 4 m/z-wide windows = 400 m/z
m/z
500
900
Scan 2
Scan 3
Notes: Bottom: DIA: 10 m/z wide windows 400 -800 10 hz scan speed
Top: MSX 5 x 2 m/z windows per scan 400 – 800 m/z 10 hz scan speed

## Slide 15: Multiplexed DIA
Scan 1
100 4 m/z-wide windows = 400 m/z
m/z
500
900
Scan 2
Scan 3
Scan 20
. . .
Notes: Bottom: DIA: 10 m/z wide windows 400 -800 10 hz scan speed
Top: MSX 5 x 2 m/z windows per scan 400 – 800 m/z 10 hz scan speed

## Slide 16: Multiplexed DIA
Scan 1
100 4 m/z-wide windows = 400 m/z
m/z
500
900
Scan 2
Scan 3
Scan 20
. . .
Notes: Bottom: DIA: 10 m/z wide windows 400 -800 10 hz scan speed
Top: MSX 5 x 2 m/z windows per scan 400 – 800 m/z 10 hz scan speed

## Slide 17: Multiplexed DIA
Scan 1
100 4 m/z-wide windows = 400 m/z
m/z
500
900
Scan 2
Scan 3
Scan 20
. . .
Notes: Bottom: DIA: 10 m/z wide windows 400 -800 10 hz scan speed
Top: MSX 5 x 2 m/z windows per scan 400 – 800 m/z 10 hz scan speed

## Slide 18: Multiplexed DIA
Scan 1
100 4 m/z-wide windows = 400 m/z
m/z
500
900
Scan 2
Scan 3
Scan 20
. . .
Scan 21
Notes: Bottom: DIA: 10 m/z wide windows 400 -800 10 hz scan speed
Top: MSX 5 x 2 m/z windows per scan 400 – 800 m/z 10 hz scan speed

## Slide 19: Multiplexed DIA
Scan 1
100 4 m/z-wide windows = 400 m/z
m/z
500
900
Scan 2
Scan 3
Scan 20
. . .
Scan 21
Notes: Bottom: DIA: 10 m/z wide windows 400 -800 10 hz scan speed
Top: MSX 5 x 2 m/z windows per scan 400 – 800 m/z 10 hz scan speed

## Slide 20: Multiplexed DIA
Scan 1
100 4 m/z-wide windows = 400 m/z
m/z
500
900
Scan 2
Scan 3
Scan 20
. . .
Scan 21
Notes: Bottom: DIA: 10 m/z wide windows 400 -800 10 hz scan speed
Top: MSX 5 x 2 m/z windows per scan 400 – 800 m/z 10 hz scan speed

## Slide 21: Multiplexed DIA
Scan 1
100 4 m/z-wide windows = 400 m/z
m/z
500
900
Scan 2
Scan 3
Scan 20
. . .
Scan 21
Notes: Bottom: DIA: 10 m/z wide windows 400 -800 10 hz scan speed
Top: MSX 5 x 2 m/z windows per scan 400 – 800 m/z 10 hz scan speed

## Slide 22: Demultiplexing
m/z
Intensity

## Slide 23: Demultiplexing
m/z
Intensity

## Slide 24: Demultiplexing
m/z
Intensity
1
7
28
81
84
Isolation Windows
Notes: Bottom: DIA: 10 m/z wide windows 400 -800 10 hz scan speed
Top: MSX 5 x 2 m/z windows per scan 400 – 800 m/z 10 hz scan speed

## Slide 25: Demultiplexing
m/z
Intensity
1
Isolation Windows
Notes: Bottom: DIA: 10 m/z wide windows 400 -800 10 hz scan speed
Top: MSX 5 x 2 m/z windows per scan 400 – 800 m/z 10 hz scan speed

## Slide 26: Demultiplexing
m/z
Intensity
1
7
28
81
84
Isolation Windows
Intensity(100) = I1 + I7 + I28 + I81 + I84
Notes: Bottom: DIA: 10 m/z wide windows 400 -800 10 hz scan speed
Top: MSX 5 x 2 m/z windows per scan 400 – 800 m/z 10 hz scan speed

## Slide 27: Demultiplexing
m/z
Intensity
3
10
74
75
92
Isolation Windows
Intensity(99) = I3 + I10 + I74 + I75 + I92
Notes: Bottom: DIA: 10 m/z wide windows 400 -800 10 hz scan speed
Top: MSX 5 x 2 m/z windows per scan 400 – 800 m/z 10 hz scan speed

## Slide 28: Demultiplexing
m/z
Intensity
Intensity(99) = I3 + I10 + I74 + I75 + I92
Intensity(100) = I1 + I7 + I28 + I81 + I84
10 Unknowns
Notes: Bottom: DIA: 10 m/z wide windows 400 -800 10 hz scan speed
Top: MSX 5 x 2 m/z windows per scan 400 – 800 m/z 10 hz scan speed

## Slide 29: Demultiplexing
m/z
Intensity
2
Knowns
10 Unknowns
Intensity(99) = I3 + I10 + I74 + I75 + I92
Intensity(100) = I1 + I7 + I28 + I81 + I84
Notes: Bottom: DIA: 10 m/z wide windows 400 -800 10 hz scan speed
Top: MSX 5 x 2 m/z windows per scan 400 – 800 m/z 10 hz scan speed

## Slide 30: Demultiplexing
…
…
Intensity(99) = I3 + I10 + I74 + I75 + I92
Intensity(100) = I1 + I7 + I28 + I81 + I84
…
…
Intensity(50) = I3 + I11 + I34 + I35 + I90
Intensity(150) = I17 + I44 + I52 + I55 + I99
100 Scans
5 Duty
Cycles
~15 seconds
100 knowns
100 unknowns
Solve by non-negative least squares optimization
Notes: Bottom: DIA: 10 m/z wide windows 400 -800 10 hz scan speed
Top: MSX 5 x 2 m/z windows per scan 400 – 800 m/z 10 hz scan speed

## Slide 31: Demultiplexing

## Slide 32: Sensitivity Similar to MS1 Quantification
Bovine proteins spiked into S. cerevisiae lysate (soluble fraction)

## Slide 33: Sensitivity Similar to MS1 Quantification
Bovine proteins spiked into S. cerevisiae lysate (soluble fraction)

## Slide 34: Conclusions
DIA data can be multiplexed by mixing precursors prior to fragment ion analysis
MSX de-multiplexing and isolation list export will be included in Skyline v1.3 (http://skyline.maccosslab.org)
A firmware patch is needed to implement this method on the Q-Exactive
  Markus Kellmann (markus.kellmann@thermofisher.com)

## Slide 35: Acknowledgments
University of Washington
MacCoss Lab
Gennifer Merrihew
Brendan MacLean
Don Marsh
Other
Ying S. Ting
Nathan Basisty
Wu Lab (University of Pittsburgh)
Nicholas Bateman
Scott Goulding
Sarah Moore
Julie Weisz
Thermo Fisher Scientific
Andreas Kuehn
Jesse Canterbury
Markus Kellmann
Vlad Zabrouskov
Funded by the
National Institutes of Health
Individual F31 fellowship -- F31 AG037265
Yeast Resource Center -- P41 GM103533